ADS1299 ECG Monitor

Complete Project Guide

NIELIT ADS1299 Board + PYNQ-Z2 FPGA

Platform: PYNQ-Z2  |  Interface: PMOD A  |  Signal: ECG/EMG/EEG/EOG

Date: March 2026

# 1. Hardware Components

## 1.1 Components Used

| Component | Details |
| --- | --- |
| FPGA Board | PYNQ-Z2 (Xilinx Zynq XC7Z020CLG400-1) |
| ADC Board | NIELIT ADS1299 Bio-Signal Acquisition Board |
| Interface | PMOD A Connector (JA) |
| PC Software | Vivado 2025.1, PuTTY, WinSCP |
| PYNQ Version | PYNQ 3.x with Python 3.10 |
| Communication | SPI Mode 0 (CPOL=0, CPHA=0), 25MHz/4 = 6.25MHz |
| Power Supply | 3.3V from PMOD A Pin 12 |

## 1.2 PMOD A Pin Connections

Connect NIELIT ADS1299 board to PYNQ-Z2 PMOD A connector exactly as follows:

| NIELIT ADS1299 Pin | PMOD A Pin | PMOD Signal Name | Zynq FPGA Pin | Function |
| --- | --- | --- | --- | --- |
| MOSI | Pin 2 (JA2) | JA2 | Y19 | SPI Data to ADS1299 |
| MISO | Pin 1 (JA1) | JA1 | Y18 | SPI Data from ADS1299 |
| SCLK | Pin 3 (JA3) | JA3 | Y16 | SPI Clock |
| CS | Pin 4 (JA4) | JA4 | Y17 | Chip Select (Active Low) |
| DRDY | Pin 7 (JA7) | JA7 | U18 | Data Ready (Active Low) |
| RST | Pin 8 (JA8) | JA8 | U19 | Reset (Active Low) |
| GND | Pin 11 | GND | - | Ground |
| 3.3V / DVDD | Pin 12 | VCC | - | 3.3V Power |
| START | Pin 12 (tied HIGH) | VCC | - | Tied to 3.3V permanently |

## 1.3 ECG Electrode Connections

| ADS1299 Input | Electrode Placement |
| --- | --- |
| IN1P (CH1+) | Left wrist inner side |
| IN1N (CH1-) | Right wrist inner side |
| IN2P (CH2+) | Left wrist outer side |
| IN2N (CH2-) | Right wrist outer side |
| BIAS | Right palm center |

# 2. Vivado FPGA Design

## 2.1 Project Settings

| Setting | Value |
| --- | --- |
| Project Path | C:/Users/mistr/ads1299_ecg/ |
| FPGA Part | xc7z020clg400-1 |
| Board | tul.com.tw:pynq-z2:part0:1.0 |
| Design Clock | 100 MHz (FCLK0) |
| Output Files | ads1299_ecg.bit + ads1299_ecg.hwh |

## 2.2 IP Blocks Used

| IP Block | Version | Base Address | Purpose |
| --- | --- | --- | --- |
| ZYNQ7 Processing System (ps7_0) | 5.5 | - | ARM CPU, AXI GP0, Clock, IRQ |
| AXI Quad SPI (axi_quad_spi_0) | 3.2 | 0x41E00000 | SPI master for ADS1299 |
| AXI GPIO (axi_gpio_0) | 2.0 | 0x41200000 | RESET output + DRDY input |
| AXI Interrupt Controller (axi_intc_0) | 4.1 | 0x41800000 | Interrupt management |
| AXI Interconnect | - | - | Connects PS to PL IPs |
| Processor System Reset | - | - | Reset logic |

## 2.3 Key IP Configuration

### AXI Quad SPI Settings:

- SCK Ratio: 4 (25MHz clock)

- FIFO Depth: 16

- Number of Slaves: 1

- SPI Memory Type: C_SPI_MEMORY = 2

- Mode: Standard SPI, Master

### AXI GPIO Settings:

- Dual Channel: enabled

- CH1: 1-bit OUTPUT (RESET pin, default HIGH = 0x00000001)

- CH2: 1-bit INPUT (DRDY pin)

- GPIO_BOARD_INTERFACE: Custom (IMPORTANT - must not be arduino_a0_a5)

- GPIO2_BOARD_INTERFACE: Custom

- Interrupt Present: enabled

## 2.4 XDC Constraints File

File: C:/Users/mistr/ads1299_ecg/pmod_a.xdc

set_property PACKAGE_PIN Y18 [get_ports spi_miso]

set_property IOSTANDARD LVCMOS33 [get_ports spi_miso]

set_property PACKAGE_PIN Y19 [get_ports spi_mosi]

set_property IOSTANDARD LVCMOS33 [get_ports spi_mosi]

set_property PACKAGE_PIN Y16 [get_ports spi_sck]

set_property IOSTANDARD LVCMOS33 [get_ports spi_sck]

set_property PACKAGE_PIN Y17 [get_ports spi_ss_n]

set_property IOSTANDARD LVCMOS33 [get_ports spi_ss_n]

set_property PACKAGE_PIN U18 [get_ports drdy_n]

set_property IOSTANDARD LVCMOS33 [get_ports drdy_n]

set_property PACKAGE_PIN U19 [get_ports reset_n]

set_property IOSTANDARD LVCMOS33 [get_ports reset_n]

# 3. PYNQ-Z2 Setup

## 3.1 Network Configuration

| Setting | Value |
| --- | --- |
| PC Ethernet IP | 192.168.2.10 |
| PC Subnet | 255.255.255.0 (/24) |
| PYNQ-Z2 IP | 192.168.2.99 |
| Connection Type | Direct Ethernet cable (PC to PYNQ) |
| SSH Port | 22 |
| Login | xilinx / xilinx |

## 3.2 PuTTY Serial Settings (for initial setup)

| Setting | Value |
| --- | --- |
| COM Port | COM6 |
| Baud Rate | 115200 |
| Data Bits | 8 |
| Stop Bits | 1 |
| Parity | None |
| Flow Control | None |

## 3.3 Commands to Set PYNQ-Z2 IP

Run in PuTTY serial terminal if IP needs to be set:

sudo ip addr flush dev eth0

sudo ip addr add 192.168.2.99/24 dev eth0

sudo ip link set eth0 up

ip addr show eth0

## 3.4 File Transfer (WinSCP)

| Setting | Value |
| --- | --- |
| Host | 192.168.2.99 |
| Port | 22 |
| Protocol | SFTP |
| Username | xilinx |
| Password | xilinx |
| Target Directory | /home/xilinx/ |

Files to transfer to /home/xilinx/:

- ads1299_ecg.bit  (from C:\Users\mistr\Downloads\)

- ads1299_ecg.hwh  (from C:\Users\mistr\Downloads\)

# 4. Running the ECG Monitor

## 4.1 How to Load Bitstream and Run

The Python script loads the bitstream automatically. Use this command every time:

sudo /usr/local/share/pynq-venv/bin/python3 /home/xilinx/ecg_monitor.py

The script will print:

- Loading overlay...  -> Overlay loaded.

- SPI CR=0x00000086 SR=0x00000025

- Resetting ADS1299...  -> Reset done.

- Device ID: 0x3E  (expect 0x3E)  <- confirms SPI working

- SPI communication OK!

- ADS1299 configured.

- Then live ECG data scrolling on screen

## 4.2 Stopping and Saving Data

Press Ctrl+C to stop. The script saves data automatically:

- CSV file saved to: /home/xilinx/ecg_YYYYMMDD_HHMMSS.csv

- Download via WinSCP (right panel shows /home/xilinx/, drag file to left panel)

# 5. Complete Python Script (ecg_monitor.py)

Save this file to /home/xilinx/ecg_monitor.py on the PYNQ-Z2:

#!/usr/bin/env python3

import os, mmap, time, csv, datetime

os.environ['XILINX_XRT'] = '/usr'

from pynq import Overlay

OVERLAY   = '/home/xilinx/ads1299_ecg.bit'

SPI_BASE  = 0x41E00000

GPIO_BASE = 0x41200000

CMD_SDATAC = 0x11; CMD_RDATAC = 0x10

CMD_START  = 0x08; CMD_STOP   = 0x0A

CMD_WREG   = 0x40; CMD_RREG   = 0x20

class MEM:

def __init__(self, base):

self.fd  = os.open('/dev/mem', os.O_RDWR | os.O_SYNC)

self.mem = mmap.mmap(self.fd, 4096, mmap.MAP_SHARED,

mmap.PROT_READ | mmap.PROT_WRITE,

offset=base)

def wr(self, off, val):

self.mem.seek(off)

self.mem.write(val.to_bytes(4, 'little'))

def rd(self, off):

self.mem.seek(off)

return int.from_bytes(self.mem.read(4), 'little')

class ADS1299:

def __init__(self):

print('Loading overlay...')

self.overlay = Overlay(OVERLAY)

print('Overlay loaded.')

self.spi  = MEM(SPI_BASE)

self.gpio = MEM(GPIO_BASE)

self._init_spi()

def _init_spi(self):

self.spi.wr(0x40, 0x0000000A)  # SRR reset

time.sleep(0.01)

self.spi.wr(0x60, 0x00000086)  # CR: SPE|Master|ManualSS

self.spi.wr(0x70, 0xFFFFFFFF)  # SSR: deassert all

time.sleep(0.01)

print(f'SPI CR=0x{self.spi.rd(0x60):08X} SR=0x{self.spi.rd(0x64):08X}')

def _transfer(self, data):

self.spi.wr(0x70, 0xFFFFFFFE)  # Assert CS

time.sleep(0.000002)

rx = []

for b in data:

self.spi.wr(0x68, b & 0xFF)

for _ in range(100000):

if self.spi.rd(0x64) & 0x04: break

time.sleep(0.000001)

rx.append(self.spi.rd(0x6C) & 0xFF)

time.sleep(0.000002)

self.spi.wr(0x70, 0xFFFFFFFF)  # Deassert CS

return rx

def wreg(self, addr, val):

self._transfer([CMD_WREG | (addr & 0x1F), 0x00, val])

time.sleep(0.00001)

def rreg(self, addr):

return self._transfer([CMD_RREG | (addr & 0x1F), 0x00, 0x00])[2]

def cmd(self, c):

self._transfer([c])

time.sleep(0.001)

def reset(self):

print('Resetting ADS1299...')

self.gpio.wr(0x00, 0)   # RESET LOW

time.sleep(0.001)

self.gpio.wr(0x00, 1)   # RESET HIGH

time.sleep(0.1)

print('Reset done.')

def init_ecg(self):

self.reset()

self.cmd(CMD_SDATAC)

dev = self.rreg(0x00)

print(f'Device ID: 0x{dev:02X}  (expect 0x3E)')

if dev != 0x3E:

print('WARNING: Check SPI wiring!')

else:

print('SPI communication OK!')

self.wreg(0x01, 0xD6)  # CONFIG1: 250 SPS

self.wreg(0x02, 0xC0)  # CONFIG2

self.wreg(0x03, 0xEC)  # CONFIG3: BIAS enabled

time.sleep(0.15)

self.wreg(0x05, 0x60)  # CH1: gain=24

self.wreg(0x06, 0x60)  # CH2: gain=24

for r in range(0x07, 0x0D):

self.wreg(r, 0x81)  # CH3-8: off

self.wreg(0x0D, 0x03)  # BIAS_SENSP: CH1+CH2

self.wreg(0x0E, 0x03)  # BIAS_SENSN: CH1+CH2

print('ADS1299 configured.')

def start(self):

self.cmd(CMD_START)

self.cmd(CMD_RDATAC)

def stop(self):

self.cmd(CMD_SDATAC)

self.cmd(CMD_STOP)

def read_frame(self):

timeout = 0

while (self.gpio.rd(0x08) & 0x01):  # Wait for DRDY LOW

timeout += 1

if timeout > 200000: return None, 0, 0

raw = self._transfer([0x00]*27)  # 3 status + 8ch x 3 bytes

def s24(b):

v = (b[0]<<16)|(b[1]<<8)|b[2]

return v - 0x1000000 if v & 0x800000 else v

return raw[:3], s24(raw[3:6]), s24(raw[6:9])

def to_uv(self, raw, gain=24, vref=4.5):

return raw * (vref/gain) / (2**23) * 1e6

def main():

print('='*55)

print(' ADS1299 ECG Monitor  |  PYNQ-Z2  |  PMOD A')

print(' Press Ctrl+C to stop and save CSV')

print('='*55)

ads = ADS1299()

ads.init_ecg()

ads.start()

ts = datetime.datetime.now().strftime('%Y%m%d_%H%M%S')

fname = f'/home/xilinx/ecg_{ts}.csv'

f = open(fname, 'w', newline='')

w = csv.writer(f)

w.writerow(['n', 't_ms', 'ch1_uv', 'ch2_uv'])

n = 0

try:

while True:

st, ch1, ch2 = ads.read_frame()

if st is None:

print('DRDY timeout - check wiring')

continue

uv1 = ads.to_uv(ch1)

uv2 = ads.to_uv(ch2)

w.writerow([n, n*4, f'{uv1:.1f}', f'{uv2:.1f}'])

if n % 25 == 0:

bar = '#' * min(50, int(abs(uv1)/200))

print(f'n={n:6d}  CH1={uv1:+8.0f}uV  CH2={uv2:+8.0f}uV  |{bar}')

n += 1

except KeyboardInterrupt:

ads.stop()

f.close()

print(f'Saved {n} samples to {fname}')

if __name__ == '__main__':

main()

# 6. ADS1299 Register Configuration

| Register | Address | Value | Description |
| --- | --- | --- | --- |
| CONFIG1 | 0x01 | 0xD6 | 250 SPS sample rate |
| CONFIG2 | 0x02 | 0xC0 | Internal reference, test signal off |
| CONFIG3 | 0x03 | 0xEC | BIAS buffer enabled, BIAS_REFINT enabled |
| CH1SET | 0x05 | 0x60 | CH1 ON, Gain=24, normal electrode input |
| CH2SET | 0x06 | 0x60 | CH2 ON, Gain=24, normal electrode input |
| CH3-8SET | 0x07-0x0C | 0x81 | CH3 to CH8 OFF (power down) |
| BIAS_SENSP | 0x0D | 0x03 | BIAS driven from CH1+CH2 positive |
| BIAS_SENSN | 0x0E | 0x03 | BIAS driven from CH1+CH2 negative |

## Data Conversion Formula

To convert raw 24-bit ADC value to microvolts:

Voltage (uV) = raw_value x (Vref / Gain) / 2^23 x 1,000,000

| Parameter | Value |
| --- | --- |
| Vref (reference voltage) | 4.5V |
| Gain | 24 |
| ADC Resolution | 24-bit (2^23 = 8,388,608) |
| Sample Rate | 250 SPS (samples per second) |
| Sample Period | 4ms per sample |

# 7. Plotting ECG Data on PC

## 7.1 Setup

Install Python libraries on your PC (run in Command Prompt):

pip install pandas matplotlib

## 7.2 Plot Script

Save as plot_ecg.py in same folder as your CSV file:

import pandas as pd

import matplotlib.pyplot as plt

df = pd.read_csv('ecg_20260310_143200.csv')  # change filename

plt.figure(figsize=(12,6))

plt.subplot(2,1,1)

plt.plot(df['t_ms'], df['ch1_uv'])

plt.title('ECG Channel 1'); plt.ylabel('uV'); plt.grid(True)

plt.subplot(2,1,2)

plt.plot(df['t_ms'], df['ch2_uv'], color='orange')

plt.title('ECG Channel 2')

plt.ylabel('uV'); plt.xlabel('Time (ms)'); plt.grid(True)

plt.tight_layout()

plt.savefig('ecg_plot.png', dpi=150)

plt.show()

## 7.3 Running the Plot Script

cd C:\Users\mistr\Downloads

python plot_ecg.py

# 8. Troubleshooting

| Problem | Cause | Solution |
| --- | --- | --- |
| Device ID: 0xFF | ADS1299 not connected or wiring wrong | Check all PMOD A connections, verify 3.3V power |
| DRDY timeout | ADS1299 not responding | Check DRDY wire on Pin 7 (JA7), check power |
| Bus error crash | PYNQ MMIO mapping bug | Use /dev/mem with mmap directly (already fixed in script) |
| No devices found | XILINX_XRT not set | Set os.environ['XILINX_XRT']='/usr' at top of script |
| ModuleNotFoundError pydantic | Wrong Python used | Use /usr/local/share/pynq-venv/bin/python3 |
| WinSCP cannot connect | IP mismatch | Set PYNQ IP to 192.168.2.99, PC to 192.168.2.10 |
| SPI CR bit 8 must be 0 | Master Transaction Inhibit | Use CR=0x086 not 0x186 (already fixed in script) |

# 9. Quick Reference Commands

## Run ECG Monitor:

sudo /usr/local/share/pynq-venv/bin/python3 /home/xilinx/ecg_monitor.py

## Check FPGA bitstream state:

cat /sys/class/fpga_manager/fpga0/state

## Check PYNQ IP address:

ip addr show eth0

## List ECG CSV files:

ls /home/xilinx/ecg_*.csv

## Edit script:

nano /home/xilinx/ecg_monitor.py

Source Code

//Copyright 1986-2022 Xilinx, Inc. All Rights Reserved.

//Copyright 2022-2025 Advanced Micro Devices, Inc. All Rights Reserved.

//--------------------------------------------------------------------------------

//Tool Version: Vivado v.2025.1 (win64) Build 6140274 Thu May 22 00:12:29 MDT 2025

//Date        : Tue Mar 10 12:28:48 2026

//Host        : LAPTOP-4HVBFQK1 running 64-bit major release  (build 9200)

//Command     : generate_target system_wrapper.bd

//Design      : system_wrapper

//Purpose     : IP block netlist

//--------------------------------------------------------------------------------

`timescale 1 ps / 1 ps

module system_wrapper

(DDR_addr,

DDR_ba,

DDR_cas_n,

DDR_ck_n,

DDR_ck_p,

DDR_cke,

DDR_cs_n,

DDR_dm,

DDR_dq,

DDR_dqs_n,

DDR_dqs_p,

DDR_odt,

DDR_ras_n,

DDR_reset_n,

DDR_we_n,

FIXED_IO_ddr_vrn,

FIXED_IO_ddr_vrp,

FIXED_IO_mio,

FIXED_IO_ps_clk,

FIXED_IO_ps_porb,

FIXED_IO_ps_srstb,

drdy_n,

reset_n,

spi_miso,

spi_mosi,

spi_sck,

spi_ss_n);

inout [14:0]DDR_addr;

inout [2:0]DDR_ba;

inout DDR_cas_n;

inout DDR_ck_n;

inout DDR_ck_p;

inout DDR_cke;

inout DDR_cs_n;

inout [3:0]DDR_dm;

inout [31:0]DDR_dq;

inout [3:0]DDR_dqs_n;

inout [3:0]DDR_dqs_p;

inout DDR_odt;

inout DDR_ras_n;

inout DDR_reset_n;

inout DDR_we_n;

inout FIXED_IO_ddr_vrn;

inout FIXED_IO_ddr_vrp;

inout [53:0]FIXED_IO_mio;

inout FIXED_IO_ps_clk;

inout FIXED_IO_ps_porb;

inout FIXED_IO_ps_srstb;

input drdy_n;

output [0:0]reset_n;

input spi_miso;

output spi_mosi;

output spi_sck;

output [0:0]spi_ss_n;

wire [14:0]DDR_addr;

wire [2:0]DDR_ba;

wire DDR_cas_n;

wire DDR_ck_n;

wire DDR_ck_p;

wire DDR_cke;

wire DDR_cs_n;

wire [3:0]DDR_dm;

wire [31:0]DDR_dq;

wire [3:0]DDR_dqs_n;

wire [3:0]DDR_dqs_p;

wire DDR_odt;

wire DDR_ras_n;

wire DDR_reset_n;

wire DDR_we_n;

wire FIXED_IO_ddr_vrn;

wire FIXED_IO_ddr_vrp;

wire [53:0]FIXED_IO_mio;

wire FIXED_IO_ps_clk;

wire FIXED_IO_ps_porb;

wire FIXED_IO_ps_srstb;

wire drdy_n;

wire [0:0]reset_n;

wire spi_miso;

wire spi_mosi;

wire spi_sck;

wire [0:0]spi_ss_n;

system system_i

(.DDR_addr(DDR_addr),

.DDR_ba(DDR_ba),

.DDR_cas_n(DDR_cas_n),

.DDR_ck_n(DDR_ck_n),

.DDR_ck_p(DDR_ck_p),

.DDR_cke(DDR_cke),

.DDR_cs_n(DDR_cs_n),

.DDR_dm(DDR_dm),

.DDR_dq(DDR_dq),

.DDR_dqs_n(DDR_dqs_n),

.DDR_dqs_p(DDR_dqs_p),

.DDR_odt(DDR_odt),

.DDR_ras_n(DDR_ras_n),

.DDR_reset_n(DDR_reset_n),

.DDR_we_n(DDR_we_n),

.FIXED_IO_ddr_vrn(FIXED_IO_ddr_vrn),

.FIXED_IO_ddr_vrp(FIXED_IO_ddr_vrp),

.FIXED_IO_mio(FIXED_IO_mio),

.FIXED_IO_ps_clk(FIXED_IO_ps_clk),

.FIXED_IO_ps_porb(FIXED_IO_ps_porb),

.FIXED_IO_ps_srstb(FIXED_IO_ps_srstb),

.drdy_n(drdy_n),

.reset_n(reset_n),

.spi_miso(spi_miso),

.spi_mosi(spi_mosi),

.spi_sck(spi_sck),

.spi_ss_n(spi_ss_n));

endmodule